## Slide 2
Time stamp information contained in each file
  The "footer" at the end of each physio log file contains several "LogStart" and "LogStop" flags. The MDH version of these flags (i.e., LogStartMDHTime and LogStopMDHTime) reflect the millisecond time stamp (since midnight) according to the same clock that is written in the DICOM header. Therefore, this clock is preferable to use to synchronize the onset of each physio log file with the onset of the DICOM collection.
  There is roughly 500ms between the onset of each log file. The exact delay between the onset of each file can be determined using the LogStartMDHTime value.
  The DICOM header time stamp (0008,0032) has a different format, but corresponds to the LogStartMDHTime flag for the last PMU log file to begin recording.
ECG

## Slide 3
Other Siemens flags'5002 ... 6002': These values mark the start (5002) and end (6002) of non-monitoring information written to the file. The information written between these two flags does not contain physiological monitoring information, and should be removed from the vector of voltage values reflecting the physio data.
'Trig': Flag written for each volume in fMRI and dMRI sequence. These are written for each volume in the preparatory scans at the beginning of the multiband sequence, as well as the volumes in the MB scan of interest.
'5000': peak in voltage change has been detected in signal
'ACQ FINISHED'
'5003':
ECG

## Slide 4
1 2 40 280 1915 1708 1531 1340 1045 ...
ECG Freq Per: 0 0
PULS Freq Per: 72 823
RESP Freq Per: 0 0
EXT Freq Per: 0 0
ECG Min Max Avg StdDiff: 0 0 0 0
PULS Min Max Avg StdDiff: 355 1646 795 5
RESP Min Max Avg StdDiff: 0 0 0 0
EXT Min Max Avg StdDiff: 0 0 0 0
NrTrig NrMP NrArr AcqWin: 0 0 0 0
LogStartMDHTime: 36632877
LogStopMDHTime: 39805825
LogStartMPCUTime: 36632400
LogStopMPCUTime: 39804637
The first five values of the first line are the parameters of the acquisition. The third of these is related to the “ticktime”, which is the time (in msec) between subsequent measurements, although the exact relationship is still being determined. Starting with the 6th value of the first line are the voltage values themselves that reflect the pulse ox. Note that the output is in voltage, not absolute SpO2 level (e.g., 98% saturation).
The sampling frequency for pulse-ox data is 50 Hz. Equivalently, a sample is acquired every 20 msecs.
Within the vector of voltage values are “trigger” events from the scanner. These are entered as 5000 (for trigger on) and 5003 (for trigger off). These values need to be stripped out of the vector. There will occasionally be extra values at the end of the voltage vector as final values in the buffer will be written to the file after the logging is stopped. This can result in the vector length being slightly longer than would be predicted from the log start and stop times described below.
The subsequent lines provide other parameters. Most relevant are the Log Start and Stop times for the MDH and MPCU. These log times indicate the time at which pulse-ox recording was started and stopped. That is, the first pulse-ox value in the first line of the data file was acquired at the LogStart time. The MPCU values provide time-stamps derived from the clock within the PMU recording system. The MDH values are time-stamps derived from the clock in the scanner, which is the same clock used to provide time-stamps for DICOM images. The MDH values are therefore preferred for synchronization of the DICOM images with the physiologic recording log.
Pulse-ox data

## Slide 5: From ICE